Drug Analysis Sytems/Instruments

Automated

Homogenous Competitive Protein Binding Immunoassays:

- Most common automated assay, for therapeutic drug monitoring (TDM) and toxicology

- Labelled Drug ( with Fluorescent, particle or enzyme) added to sample with compete for binding (to limited quantity of antibody) sites with drug in sample.

- In Fluorscently labeled assays, plane polarized light is passed through sample, in Particle labeled, insoluble aggregates form and produce light scatter (measured at 340 nm), and in Enzyme mediated labeling, emzymes that bind antibodies are inactivated

- In Enzyme mediated assay, any drug linked enzyme is inactivated on binding with antibody. Unbound enzyme activity proportional to drug concentration

- Drug concentration in sample inversely proportional to fluorescence and particle aggrevation, and in enzyme mediated, drug concentration in sample is proportional to the products of enzyme mediated reaction (usually NAD, NADH).

Alcohol Dehydrogenase Automated Test for Ethanol

- Ethanol + NAD --------- Acetaldehyde + NADH (absorbance detected at 340 nm) ….absorbance is proportional to Ethanol Concentration

Automated Lipid Measurement Methods

- ex. Roche Modular System …..Cholesterol and TG reagents are used to quantitate concentrations using a timed enzymatic end point method with color measurements at 520 or 560.

- Use Hydrogen peroxide (H202) a lot to react with dyes creating color to quantitate sample concentrations

Cholesterol – Cholesterol Ester(cholesterol esterase)Cholesterol(cholesterol oxidase) Cholestene-3-one + H202 + 4-AAP + Phenol (peroxidase) = quinonemine (measured at 520m)

Triglyceride – TG(lipase) Gylcerol + O2  (Glycerol Kinase) G3P  (G3P oxidase) Dihydroxyacetone + H2O2 H2O2 + 4-AAP  (peroxidase) Quinonemine (measured at 520 nm)

HDL-Cholesterol – detergent that solubilizes only HDL, releasing HDL cholesterol for subsequent reactions. Same detergents also absorbs VLDL LDL and chylomicrons, preventing their rxns.

Cholesteryl Ester  (cholesterol esterase + cholesterol oxidase) Delta 4 Cholestenon + H2O2 H2O2 + 4 AAP  (peroxidase) Color product (560 nm)

LDL – Cholesterol – Most commonly we use a calculation…… TC = VLDL-C + LDL-C + HDL-C

LDL-C = TC – (TG/2.2 + HDL-C) …..cannot be used with very high TG values (skews VLDL assumption)

Chromatography Resolution : GC > HPLC > TLC

A method to separate components of a sample between two phases, a stationary one(solid or liquid supported by a solid or gel) and a mobile one (liquid or Gas)

Gas Chromatography

- Glass tube supports the silicone polymer stationary phase, Mobile phase is the carrier gas.

- Separation based on the vapourization of sample and the presence polar functional groups.

- Used to analyze drugs that can be vapourize, and are thermally stable

- Higher resolution than both HPLC and TLC – drugs identified by retention time (attraction to the stationary phase, polar functional froups)

High Pressure Liquid Chromatography – Mainly for TDM

Minute silica particles coated in hydrophobic fuzz are the stationary phase, supported by a stainless steel tube, and Solvent Mixture of sample is the mobile phase. Drugs identified by rentention time, based on presence of polar functional groups (polarity). Can be used for thermally unstable drugs and compounds, done under 100C.

Thin Layer Chromatography

Stationary phase is dried silica, support on glass or plastic plate…mobile phase is a solvent mixture…..compunds separated based on polarity and migration on the stationary phase.

Gas Chomatography/Mass Spectrometry (GC/MS)

- Sample components separated by GC and analyzed by Mass Spectrometer, which fragments compounds after they exit GC….MS monitors scan molecular mass range in which molecules and their fragments are expected to occur.

- This creates a total ion chromatograph (from GC) and a mass spectrum (fragmentation pattern/finger print)

- Identification by rentition time and fragmentation pattern

- This is the gold standard of drug confirmation, and is legally accepted

HPLC/MS/MS

- HPLC combined with 2 quadrupole mass spectrometers. Combination with the mass specs gives the instrument more analytical sensitivity and freedom from interferences

- A special electrospray interface between HPLC and MS removes solvent (uncharged) while fragmenting and ionizing the analyte molecules before they enter the first MS.

- Mass specs can be set to analyze whole mass range of sample, or to look for specific mass ions (selective ion monitoring)

- Typically, only one ion fragment enters the first quadrupole, and collides with nitrogen in the second collision cell, creating smaller ions…..two of these smaller ions enter the third quadrupole, one to give the identity of the parent compound (qualifier ion), and the other target ion to quantitate the amount in the sample. After the 3rd quadropole they are measure by the detector, an electron mutilplier device. Called Multiple Reaction monitoring…..some original signal is lost, but background is eliminated, boosting the signal/noise ration. Sensitivity and specificity are greatly increase.

- Used for toxicology, identification is by retention time and mass spectrum (ion fragments)

- Used for measuring extremely low concentrations of immunosuppressive drugs, vitamin D and hormones